System Settings – Network and Radio

PG4 node: ____ BP2 node: ____ Script slot: ____

Radio AP Address: ____ PM Address: ____ Channel: ____

### Stimulation Cheat Sheet

% typical stimulation pathway nnp.networkOn; nnp.setSync(25); % 40 Hz, SYNC ms = 1000/Hz nnp.enterTestStim; nnp.write(node, '3212', channel, [PW_us, PA_mA*10]); % set stim on at node/channel nnp.write(node, '3212', channel, [0, 0]); % set stim off at node/channel nnp.enterWaiting; nnp.networkOff

| Parameter | Range | Note |
| --- | --- | --- |
| Pulse Width (PW) | 1–255 µs | 1 µs resolution |
| Pulse Amplitude (PA) | 0–20 mA | written as 0–200, i.e. mA × 10 |
| Inter-phase interval | 5–100 µs | default 50 µs |
| SYNC / channel period | 10–65535 ms | default 80 ms = 12.5 Hz |

### Troubleshooting

- Error on NNP-API function or app call?

  - Make sure you have the correct directories and subdirectories in your path

- Modules not appearing on network?

  - Check that switches are pointing away from BOOT and RESET

  - Check jumpers are placed on network pins

  - Make sure modules are not mapped to the same node number

- Wireless Link not communicating with Power Module

  - Check power and USB connection

  - Increase the transmit power and place the antennae closer.

  - Refresh the serial port

- Always track changes to radio settings!

- Anything else…

  - Put a hand up and we’ll come help!

### Engage with the Community

Sign up for the Community Forum and follow on LinkedIn through landing.cosmiic.org

### Post-Session Feedback

Five minute survey after the session so we can improve the learning experience